	Objective
	Magnification
	1 division
	10 divisions
	100 divisions

	3.2x
	32x
	32 microns
	320 microns
	3.2 mm

	10x
	100x
	10
	100
	1 mm

	16x
	160x
	6.3
	63
	0.63 mm

	40x
	400x
	2.5
	25
	0.25 mm

	100x
	1000x
	1.0
	10
	0.1 mm

	NB:  The 3.2x is rarely used [the Wild M5 would be better.
	The 100x objective requires the use of immersion oil.
	
	
	

	
	
	
	
	


My Zeiss microscope calibration 

This is for use with 10x eypiece with graticule [calibrated with a 1mm scale]

